## Slide 1
Real-Time Mode
Click to Start
Blue: heart rate, respiration, or temperature exceed/below the limit
Yellow: battery status, lead off, connection failure
Red: asystole, VT, VF
Patient Name
One-hour status bar
Patient ID
Time = 1 sec
Gain
Click to adjust display gain
Click to select (a), (b), (c), or (d)
(a) auto
(b) 1mV (default)
(c) 2 mV
(d) 0.5 mV
Current time
Sign in
Print all patients’ latest 8 sec. waveforms
User manual
Home
Setting
Retrieve
Exit
Asystole (Red)
Below the alarm limit (Blue)
Lead off (Yellow)
Connection failure
(Yellow)
HR
Temp.
RR
SpO2/ BP
(External)
Click
for setting
Please recharge within 8 hours.
Battery Level
H
L
Location
(while roaming)

## Slide 2
Retriever Mode
(review data history)
Start time of trend
Start time of one-hour waveforms
Time & vital signs of standard waveforms
Click to display selected vital sign in trend
End time of standard waveforms (15 sec)
Move the green block to display standard waveforms
(white block: print region)
End time of one-hour waveforms
Move the green block to display one-hour waveforms
Trend duration zoom in/out (1/8/24/120 hr)
End time of trend
Move/click the full time scale to select time region of trend
Print standard waveforms
Print statistical report

## Slide 3
Admission
Click to set the upper/ lower limit
Click on a digit
Click increase/ decrease
Click on confirm (V)/ cancel (X)
Snap two electrodes and stick to patient
Click to confirm
10 sec testing signals
Discharge
Click on the ward/ bed # to continue
Click on the discharge icon
Click to confirm
Take off the transducer and recharge
Discharge completed
Retry
Cancel
Click on any blank region to begin
Select ward/bed #
Enter patient ID (*required) and name
Click to admit
*Required
Press “heart” to power on
Notes: To Check:
15sec測試訊號？

## Slide 4
Transfer
Transducer Swap
Retrieve (Discharged)
Click on the ward/ bed #
Click on the ward/ bed #
Select a new ward/ bed #
Click to confirm
Click to transfer
Click retrieve icon (menu bar) to load retriever
Click to retrieve
Choose one of the options below:
Input patient ID.
Set possible admission time range.
Input keyword of patient ID.
Input keyword of patient ID and set possible admission time range.
Click Search button to search.
Monitor with the new transducer
Press the” heart” on the new transducer
Click on the ward/ bed # to continue
Click to swap transducer
Click to confirm
Recharge the used transducer
Click to swap
Notes: To Check:
15sec測試訊號？